| URS for HPHV Steam Sterilizer-QC |  |
| --- | --- |
| Equipment | HPHV Steam Sterilizer |
| URS number |  |
| Version No. | 01 |
| Effective Date | Date of approval |

This document and its contents are the confidential information of ABC Private Limited, Unit-2,

No part of this document may be copied, reproduced, modified, or distributed to any third party, in any form or by any means, or stored in a database or retrieval system, without prior written permission of ABC Private Limited.

Table of Contents

1.	Overview and Description	3

2.	Facility Requirement	3

3.	Technical Data and System Specification	3

4.	Automation and Instrumentation	6

5.	Utility Requirements	6

6.	Other Requirements	7

6.1.	Cleaning Requirements	7

6.2	Sterilization Requirements	7

6.3	Calibration services	7

6.4	Preventive maintenance and Breakdown maintenance services	7

6.5	Safety Requirements	7

6.6	Any other details	7

7	Revision History	7

8	Approval Details	8

# Overview and Description

A HPHV Steam sterilizer or autoclave is a pressure chamber used to carry out industrial processes requiring elevated temperature and pressure different from ambient air pressure. Many autoclaves are used to sterilize equipment and supplies by subjecting them to high pressure saturated steam at 121°C to 124°C for around 15–20 minutes depending on the size of the load and the contents.

This document defines only the user requirements for the HPHV Steam Sterilizer, its utility and ancillary equipment that can be accommodated in the Quality Control of a cGMP facility.

# Facility Requirement

The HPHV Steam Sterilizer is intended to be located in the “Autoclave Loading area”, within the {{f:facility_code}} facility. The room layout is enclosed along with the facility URS. The vendor is expected to locate the system in this layout, and clarify all required space for operation and maintenance, including any technical areas.

Area -  2.5 × 3.0 meters approximately.

Operating Environment – Loading side should be ISO 8 and Unloading side should be ISO 7.

# Technical Data and System Specification

System number: S-56

| - Sl. No. | - Specification | - Requirement |
| --- | --- | --- |
|  | - Operating Principle | - Should use saturated steam under constant pressure as sterilizing agent to maintain sterilization temperature. |
|  | - Model | - Horizontal, HPHV Steam sterilizer |
|  | - Capacity | - 320 L approximately |
|  | - Mounting | - Floor mounting |
|  | - Sterilization cycle requirements | - Both HPHV cycle and Standard cycle should be there |
|  | - Chamber | - Rectangular chamber with 6mm thickness |
|  | - Shelves | - 2 No’s of SS316L shelves with perforations are required. |
|  | - Sensors | - Should be RTD PT-100 - 4 Sensors should be in Chamber - 1 Sensor should be in Drain |
|  | - Trolley | - Loading and Unloading trolleys are required |
|  | - Instrument Temperature range | - 115°C to 150°C |
|  | - Working Temperature range | - 115°C to 134°C |
|  | - Working pressure | - 1.2 to 2.2 kg/cm² |
|  | - Pressure gauge | - 0 to 6 kg/cm² for jacket |
|  | - Compound gauge | - 1 to 6 kg/cm² for chamber |
|  | - Safety valve range | - 0.0 to 3.0 kg/cm² |
|  | - Emergency stop | - Required |
|  | - External Dimensions | - Should be proportionate to internal dimension |
|  | - Internal Dimension (HxWxD) mm | - Vendor to specify |
|  | - Material of construction | - Outer body - Stainless steel 304 - Inner chamber - Stainless steel 316 L - Door - Stainless steel 316 L with Ra value 0.8 micron or better - Jacket - Stainless steel 304 - Cabinet enclosure - Stainless steel 304 |
|  | - Door thickness | - 50-60mm approximately |
|  | - Double door/single door | - Fully automated, pneumatic horizontal sliding double doors are required |
|  | - Door safety features required | - Door interlock to prevent simultaneous opening of both doors. - Process-Lock to prevent opening of the door during the process - Process-Lock to prevent the opening of the non-sterile door after successful completion - of the cycle - Door obstruction safety - Cross contamination seal |
|  | - Door gasket | - Food grade Silicon gasket required |
|  | - Door operating cylinder | - Required |
|  | Insulation | - Glass wool or mineral wool for both chamber and jacket |
|  | Jacket thickness | - 5-6mm approximately |
|  | Piping | - All piping in contact with the chamber should be fabricated with SS 316 L with argon welding. |
|  | Connection to Drain piping | - Required |
|  | Sterilizer operation | - Automatic & manual |
|  | Process Management System (Controller) | - Should be with PLC/DCS - 3 level password protection for process security - RS 232 port for data communication - In-built Real Time Clock with date and time function - Battery backup for automatic restart in case of power failure - Facility to force all field devices via keys on the HMI for manual operation or maintenancell these features should be there in the controller. |
|  | Strip chart recorder | - Six-Channel Strip Chart Recorder for recording temperature at 5 points and pressure at 1 point is required. |
|  | Battery backup | - Should provide battery backup for automatic restart in case of power failure. |
|  | Steam Generator | - Required |
|  | Temperature and Pressure sensors | - Should have adequate sensors to monitor and control chamber temperature and pressure. |
|  | Validation ports | - Validation ports with tri-clamp connections are required for periodic validation |

# Automation and Instrumentation

  - Application Software

21 CFR Part 11 compliance software.

  - Data and Security Measures

| - Sl. No. | - Specification | - Data and security requirement |
| --- | --- | --- |
|  | - Access levels | - Manager - Admin level - Supervisor level - Operator level |
|  | - Frequency of data collection | - Every sterilization cycle |
|  | - Mode of data collection - [Hard copy/soft copy] | - Hard copy/soft copy required |
|  | - Data retention time | - Required |
|  | - Interface with other equipment’s[communication port/modem connections] | - PC & printer |
|  | - Software Version | - Vendor to specify |
|  | - IT requirements | - PC, LAN connection & 21 CFR part 11 compliance software |

# Utility Requirements

The following utilities should be available at the facility and the same is required for the functioning of the HPHV Steam Sterilizer.

| Utility | Requirements |
| --- | --- |
| Power requirements | 400-440±10% V, 3 phase, 50±5% Hz |
| UPS Requirement | UPS power backup is required in case of any power failure |
| Water & steam supply | Purified water, Cooling water and Pure steam supply required. |
| Compressed air requirement. | 0.2 m3/hr. during peak hours at 6-7 Kg/cm2 pressure. |

# Other Requirements

## Cleaning Requirements

Cleaning should be performed as per the guidelines given in the equipment user manual.

## Sterilization Requirements

Sterilization should be both automatic & manual.

## Calibration services

Calibration of all critical instruments to be performed. Calibration certificates to be provided or all instruments. All calibration certificates to have traceability to national and international standards.

## Preventive maintenance and Breakdown maintenance services

Support on Preventive maintenance and Breakdown maintenance has to be provided based on In-house procedure.

## Safety Requirements

- All the wiring connection should have proper insulation.

- Door locking facility

- Low water protection system

- Pressure cut off facility

- Emergency stop

- Air pressure low alarm

## Any other details

| - 1 | Warranty | - Minimum 1 Year |
| --- | --- | --- |
| - 2 | Training | - To the User group and maintenance department. |
| - 3 | Other Supports | - Installation at Site |
|  |  | - List of accessories and spares |
|  |  | - Support for calibration or re-qualification |
|  |  | - Should perform and document the Qualification like DQ/IQ/OQ/PQ at the site should provide all the qualification & test certificates. |
|  |  | - All certificate of compliance need to be provided. |

# Revision History

| Version Number | Revision summary | Effective Date |
| --- | --- | --- |
| 01 | Newly Introduced | Date of Approval |

# Approval Details

By signing this document all the concern agrees to the content of the documents.

| - Department | - Name | - Sign | - Date |
| --- | --- | --- | --- |
| - Prepared by |  |  |  |
| QC |  |  |  |
| - Reviewed by |  |  |  |
| QC |  |  |  |
| ED |  |  |  |
| EHS |  |  |  |
| IT |  |  |  |
| QA |  |  |  |
| - Approved by |  |  |  |
| QA |  |  |  |